Function Specifications

for

the Model 15 Amplifier System

(Second Draft)

Author: Aarti Goge

Supervisor: Prof. Adrian Chan

Date of Submission: June 23, 2003

1.0 Introduction:

The Model 15 amplifier system is a programmable, modular, multi-channel signal conditioning system.  It provides an interface between various biomedical signal sources and data acquisition systems. The signal sources can be preparations having electrodes, transducers, and other special sensors required for the desired protocol. The data acquisition system can range from oscilloscopes to sophisticated computer-based systems utilizing Grass PolyVIEW, Gamma and other software packages. The system has patient isolation so that it can be used for human recording applications.

The Model 15 has eight amplifier plug-in slots. The current system has two 15A54 Quad, wide-band, high-gain, programmable AC amplifiers in slots 1 and 2. All other slots are empty. Each amplifier module has four discrete channels of amplification. The current system with two quad amplifiers provides eight channels.  Most of the system functions can be controlled via RS-232 serial port.

This is a function specification document, explaining various command functions for controlling the Model 15 amplifier system behavior. It provides the basic building blocks for programmatically controlling this system in the Matlab environment. The following section provides a brief description of all the command functions classified as general, system and amplifier commands.

2.0 Commands:

Each subsection begins with a block diagram illustrating the co-relation within all the command functions followed by a brief description of the individual functions.

The function format is as follows:

function outputs = M15_FunctionName(inputs)

function purpose

input arguments

function algorithm

output arguments

All the system and amplifier commands sent to the Model 15 are short ASCII strings beginning with the escape character and ending with a 2-byte checksum and carriage return:

command = [esc, address, command, amplifiernum, parameter1, parameter2, checksum, cr]

The parts of the commands sequence that are in italics are considered “optional” parameters, since not all of the commands require these parameters.

[address]: This is the address of the Model 15 system, from 1 to 8. This must match the address setting that is selected with the ID Number dip switch in the upper left hand corner of the Controller module. The address must be a single byte, ASCII-encoded decimal. The default factory setting is 1.

[command]: The command is 1-byte, uppercase ASCII letter representing the desired command.

[amplifiernum]: The amplifiernum is a 2-byte sequence from 0 to 32 in ASCII-encoded hexadecimal (“00” to “20”), specifies which amplifier the command will affect. To apply the command to all amplifiers for the given Model 15 system, specify “00” for the amplifiernum parameter.

[parameter1][parameter2]: parameter1 and parameter2 are only required for some commands and are explained in the command description that follows.

[checksum]: The checksum is a 2-byte, ASCII-encoded hexadecimal result of adding all of the bytes sent, including the “escape” character, but not including the checksum or the final carriage return. The result is truncated to 2 bytes, if necessary. A checksum of 243 decimal (F3 hex) must be transmitted as the 2 characters “F” followed by “3”.

2.1 General Commands:

Block diagram explaining co-relation within all the general command functions.

M15_InitializeSerialPort

M15_DestroySerialPort

M15_SerialRead

M15_SerialWrite(command)

M15_CheckSum(command)

M15_StatusCode(response)

M15_VerifyCode(code,errormsg)

The following functions are used by both system and amplifier command functions:

function M15_InitializeSerialPort

M15_InitializeSerialPort creates a serial port object and connects it to the device.

input = none

global variables = M15_SERIALPORT, the serial port name (string)

= M15_OBJ, the serial port object

This function makes use of built-in Matlab functions as follows:

To create a serial port object, we use:

M15_OBJ = serial(M15_SERIALPORT)

To connect the serial port object to the device, we use:

fopen(M15_OBJ)

output = none

function M15_DestroySerialPort

M15_DestroySerialPort destroys a serial port object.

input = none

global variable = M15_OBJ, the serial port object

To disconnect the serial port object from the device, we use the built-in Matlab function as follows:

fclose(M15_OBJ)

output = none

function response = M15_SerialRead

M15_SerialRead reads data from the device.

input = none

global variable = M15_OBJ, the serial port object

To read data from the device, it makes use of built-in Matlab function as follows:

response = fscanf(M15_OBJ)

output = response, reply from the device (string)

function M15_SerialWrite(command)

M15_SerialWrite adds escape character and address at the beginning of the command.

It adds checksum and carriage return at the end of the command and writes the command

to the device.

input = command (string)

constants = M15_ESC (27 decimal)

= M15_CR (13 decimal)

global variable = M15_OBJ, the serial port object

= M15_ADDRESS, the address of the Model 15 system (decimal)

command = [M15_ESC, M15_ADDRESS, command]

command = [command, M15_CheckSum(command), M15_CR], ASCII string ending with the ASCII “carriage return” character

To write data to the device, it makes use of built-in Matlab function as follows:

fprintf(M15_OBJ, command)

output = none

function checksum = M15_CheckSum(command)

M15_CheckSum adds all of the bytes to be sent, including the escape character, but not including the checksum or final carriage return. The result is a 2-byte, ASCII encoded hexadecimal and is truncated to 2 bytes, if necessary. A checksum of 243 decimal (F3 hex) must be transmitted as 2 characters “F” followed by “3”.

input = command, ASCII string beginning with the ASCII “escape” character

This function makes use of built-in Matlab functions as follows:

To sum all of the bytes to be sent and to write formatted data to a string:

checksum = sprintf('%X', sum(abs(command)));

To get the length of the checksum vector:

N = length(checksum)

(N should be 2 bytes long)

output = checksum (2 characters)

function [code, errormsg] = M15_StatusCode(response)

The Model 15 responds to most commands with “OK<cr>”, if the command is accepted. Otherwise, a 2-byte error code “XX<cr>” is sent.

M15_StatusCode resolves the response into appropriate code and error message.

input = response (string)

global variable: M15_OK

constants: M15_OK = 0

M15_CM = 1

M15_CK = 2

M15_CH = 3

M15_VU = 4

To compare the first n characters of two strings, it makes use of built-in Matlab function as follows:

k = strncmp(‘str1’,‘str2’,n)

Returns logical true (1) if the first n characters of the strings str1 and str2 are the same, and returns logical false (0) otherwise.

‘str1’ varies with ‘OK’, ‘CM’, ‘CK’, ‘CH’, ‘VU’ and compares for equality the first two characters of the response as ‘str2’.

output = code (integer)

= errormsg (string)

Table 2.1.1 Status Codes

| response | code | errormsg |
| --- | --- | --- |
| OK<cr> | M15_OK | command accepted/executed |
| CM<cr> | M15_CM | command/data error |
| CK<cr> | M15_CK | checksum error |
| CH<cr> | M15_CH | invalid channel number |
| VU<cr> | M15_VU | invalid setting or value |

function M15_VerifyCode(code, errormsg)

M15_VerifyCode outputs an error message if the code is not equal to M15_OK.

input = code (integer)

= errormsg (string)

global variable: M15_OK

if (code ~= M15_OK)

To output the errormsg, it makes use of built-in Matlab function as follows:

fprintf(errormsg);

end

output = none

2.2 System Commands:

Block diagram explaining co-relation within all the system command functions.

M15_Initialize

M15_SerialRead

M15_QueryStatus

M15_DcCalEnable(setting)

M15_CalSettings(volfreq, setting)

M15_SetMode(setting)

M15_TraceRestore(setting)

M15_WhoYouAre

M15_SerialWrite(command)

M15_StatusCode(response)

M15_CheckSum(command)

M15_VerifyCode(code, errormsg)

M15_QueryID

function M15_WhoYouAre

This command must be sent first. It tells the Model 15 what type of amplifier module is in each slot. Without this other commands will not be sent properly. This should be sent every time a new connection is established with the Model 15.

M15_WhoYouAre creates and writes a command to the system. The system responds with appropriate message and outputs an error message if any.

input = none

command = [‘F’, WHO 1, WHO 2, WHO 3, WHO 4, WHO 5, WHO 6, WHO 7, WHO 8]

For the current system with two 15A54 amplifiers in the first two slots and remaining six slots empty, the command is of the form:

command = [‘F00999999’]

(Please refer Table 2.2.1 WHO X)

M15_SerialWrite(command)

response = M15_SerialRead

[code, errormsg] = M15_StatusCode(response)

M15_VerifyCode(code, errormsg)

output = none

Table 2.2.1 WHO X

| setting | hexadecimal | char |
| --- | --- | --- |
| 15A54 or 15A94 Amplifier | 30 | 0 |
| 15A12 Amplifier | 31 | 1 |
| Empty slot, 15A04 or 15A02 | 39 | 9 |

function M15_Initialize

This command initializes all amplifiers to the non-volatile default settings and clears any pending errors.

M15_Initialize creates and writes a command to the system. The system responds with appropriate  message and outputs an error message if any.

input = none

command = [‘I’]

M15_SerialWrite(command)

response = M15_SerialRead

[code, errormsg] = M15_StatusCode(response)

M15_VerifyCode(code, errormsg)

output = none

function firmwarenum = M15_QueryID

This command responds with the Model 15 firmware revision number.

M15_QueryID creates and writes a command to the system. The system responds with appropriate error message followed by a firmware revision number. The function outputs the firmware revision number.

input = none

command = [‘U’]

M15_SerialWrite(command)

firmwarenum = M15_SerialRead

(firmwarenum = “OK<cr>GRASS Model15 Rev.XX.XX”)

output = “GRASS Model 15 Rev.XX.XX”

function M15_QueryStatus

This command checks the status of the Model 15 and responds with either “OK” or the last encountered error code “XX”.

M15_QueryStatus creates and writes a command to the system. The system responds with appropriate message and the function outputs an error message if any.

input = none

command = [‘E’]

M15_SerialWrite(command)

response = M15_SerialRead

[code, errormsg] = M15_StatusCode(response)

M15_VerifyCode(code, errormsg)

output = none

function M15_SetMode(setting)

This command sets the Model 15 into either the Calibration or Use mode (to switch all of the amplifier inputs to either the calibrator circuit or the electrode inputs).

M15_SetMode creates and writes a command to the system. The system responds with appropriate message and the function outputs an error message if any.

input = setting (string ‘use’ or ‘cal’)

command = [‘C’, parameter1]

M15_SerialWrite(command)

response = M15_SerialRead

[code, errormsg] = M15_StatusCode(response)

M15_VerifyCode(code, errormsg)

output = none

Parameter1 should be sent in ASCII-encoded hexadecimal, representing the following settings:

Table 2.2.2 Mode Setting

| setting | parameter1 |
| --- | --- |
| use | 0 |
| cal | 1 |

function M15_CalSettings(volfreq, setting)

This command sets the calibrator frequency and amplitude. One of the two inputs designates whether the command is for the frequency or amplitude setting. The other input designates the setting itself.

M15_CalSettings creates and writes a command to the system. The system responds with appropriate message and the function outputs an error message if any.

input = volfreq (string ‘voltage’ or ‘frequency’)

= setting (decimal)

command = [‘K’, parameter1, parameter2]

M15_SerialWrite(command)

response = M15_SerialRead

[code, errormsg] = M15_StatusCode(response)

M15_VerifyCode(code, errormsg)

output = none

Parameter1 and parameter2 should be sent in ASCII-encoded hexadecimal, representing the following settings:

Table 2.2.3 Calibrator Setting

| volfreq | parameter1 |
| --- | --- |
| voltage | A |
| frequency | F |

| setting (in V) | parameter2 (Voltage) |
| --- | --- |
| 5 | 0 |
| 10 | 1 |
| 20 | 2 |
| 50 | 3 |
| 100 | 4 |
| 200 | 5 |
| 500 | 6 |
| 1000 | 7 |

| setting (in Hz) | parameter2 (Frequency) |
| --- | --- |
| DC | 0 |
| 0.3 | 1 |
| 1 | 2 |
| 3 | 3 |
| 10 | 4 |
| 30 | 5 |
| 100 | 6 |
| 300 | 7 |
| 1000 | 8 |

function M15_DcCalEnable(setting)

This command applies or removes a DC calibration signal, when the system is in the Calibration mode and the frequency is set to 0 for DC.

M15_DcCalEnable creates and writes a command to the system. The system responds with appropriate message and the function outputs an error message if any.

input = setting (string ‘off’ or ‘on’)

command =  [‘D’, parameter1]

M15_SerialWrite(command)

response = M15_SerialRead

[code, errormsg] = M15_StatusCode(response)

M15_VerifyCode(code,errormsg)

output = none

Parameter1 should be sent in ASCII-encoded hexadecimal, representing the following settings:

Table 2.2.4 DC Calibrator Enable Setting

| setting | parameter1 |
| --- | --- |
| off | 0 |
| on | 1 |

function M15_TraceRestore(setting)

This command applies or removes the trace restore function (clamp amplifiers).

M15_TraceRestore creates and writes a command to the system. The system responds with appropriate message and the function outputs an error message if any.

input = setting (string ‘off’ or ‘on’)

command = [‘A’, parameter1]

M15_SerialWrite(command)

response = M15_SerialRead

[code, errormsg] = M15_StatusCode(response)

M15_VerifyCode(code,errormsg)

output = none

Parameter1 should be sent in ASCII-encoded hexadecimal, representing the following settings:

Table 2.2.5 Trace Restore Setting

| setting | parameter1 |
| --- | --- |
| off | 0 |
| on | 1 |

2.3 Amplifier Commands:

Block diagram explaining co-relation within all the amplifier command functions.

M15_SetLineFilter(setting)

M15_ElectrodeTest(setting)

M15_SetHighFilter(setting, amplifiernum)

M15_SaveDefaults

M15_SetGain(setting, amplifiernum)

M15_SetGainRange(setting, amplifiernum)

M15_SetLowFilter(setting, amplifiernum)

M15_SerialWrite(command)

M15_CheckSum(command)

M15_QuerySettings(amplifiernum)

M15_SerialRead

M15_StatusCode

M15_VerifyCode(code,errormsg)

function [highfilter, linefilter, gainrange, gain, lowfilter] = M15_QuerySettings(amplifiernum)

This command reads back the amplifier’s settings specified by the amplifiernum as an input. The response will provide basic characteristics of the desired amplifier(high filter, line filter, gain range, gain, low filter).

M15_QuerySettings creates and writes a command to the system. The system responds with appropriate message and the setting of the desired amplifier. The function outputs an error message if any and the setting.

input = amplifiernum (2-byte sequence from 0 to 32 in ASCII encoded hexadecimal, “00” to “20”)

command = [‘Q’, amplifiernum]

M15_SerialWrite(command)

response = M15_SerialRead

[code, errormsg] = M15_StatusCode(response)

M15_VerifyCode(code,errormsg)

response = M15_SerialRead

(response = [esc, address, S, amplifiernum, High Filter, Line Filter, Gain Range, Gain, Low Filter, checksum, cr])

output = [highfilter, linefilter, gainrange, gain, lowfilter] (Refer Tables: 2.3.3, 2.3.2, 2.3.5, 2.3.6, 2.3.4)

function M15_ElectrodeTest(setting)

This command initiates an electrode test sequence. The test can be turned off or on.

M15_ElectrodeTest creates and writes a command to the system. The system responds with appropriate message and the function outputs an error message if any.

input = setting (string ‘off’ or ‘on’)

command = [‘T’, parameter1]

M15_SerialWrite(command)

response = M15_SerialRead

[code, errormsg] = M15_StatusCode(response)

M15_VerifyCode(code,errormsg)

output = none

Parameter1 should be sent in ASCII-encoded hexadecimal, representing the following settings:

Table 2.3.1 Electrode Test Setting

| setting | parameter1 |
| --- | --- |
| off | 0 |
| on | 1 |

function M15_SetLineFilter(setting, amplifiernum)

This command changes the line frequency filter setting for the designated amplifier to either off or on.

M15_SetLineFilter creates and writes a command to the system. The system responds with appropriate message and the function outputs an error message if any.

input = setting (string ‘off’ or ‘on’)

= amplifiernum (2-byte sequence from 0 to 32 in ASCII encoded hexadecimal, “00” to “20”)

command = [‘N’, amplifiernum, parameter1]

M15_SerialWrite(command)

response = M15_SerialRead

[code, errormsg] = M15_StatusCode(response)

M15_VerifyCode(code,errormsg)

output = none

Parameter1 should be sent in ASCII-encoded hexadecimal, representing the following settings:

Table 2.3.2 Line Filter Setting

| setting | parameter1 |
| --- | --- |
| off | 0 |
| on | 1 |

function M15_SetHighFilter(setting, amplifiernum)

This command changes the high-filter setting for the designated amplifier to the setting represented by the input.

M15_SetHighFilter creates and writes a command to the system. The system responds with appropriate message and the function outputs an error message if any.

input = setting (decimal)

= amplifiernum (2-byte sequence from 0 to 32 in ASCII encoded hexadecimal, “00” to “20”)

command = [‘H’, amplifiernum, parameter1]

M15_SerialWrite(command)

response = M15_SerialRead

[code, errormsg] = M15_StatusCode(response)

M15_VerifyCode(code,errormsg)

output = none

Parameter1 should be sent in ASCII-encoded decimal, representing the following settings:

Table 2.3.3 High Filter Setting

| 15A54 setting | setting | parameter1 |
| --- | --- | --- |
| 30 Hz | 30 | 0 |
| 100 Hz | 100 | 1 |
| 300 Hz | 300 | 2 |
| 1 kHz | 1000 | 3 |
| 3 kHz | 3000 | 4 |
| 6 kHz | 6000 | 5 |

function M15_SetLowFilter(setting, amplifiernum)

This command changes the low-filter setting for the designated amplifier to the setting represented by the input.

M15_SetLowFilter creates and writes a command to the system. The system responds with appropriate message and the function outputs an error message if any.

input = setting (decimal)

= amplifiernum (2-byte sequence from 0 to 32 in ASCII encoded hexadecimal, “00” to “20”)

command = [‘L’, amplifiernum, parameter1]

M15_SerialWrite(command)

response = M15_SerialRead

[code, errormsg] = M15_StatusCode(response)

M15_VerifyCode(code,errormsg)

output = none

Parameter1 should be sent in ASCII-encoded decimal, representing the following settings:

Table 2.3.4 Low Filter Setting

| setting (in Hz) | parameter1 |
| --- | --- |
| 0.01 | 0 |
| 0.1 | 1 |
| 0.3 | 2 |
| 1.0 | 3 |
| 3.0 | 4 |
| 10 | 5 |
| 30 | 6 |
| 100 | 7 |

function M15_SetGainRange(setting, amplifiernum)

This command changes the gain range setting for the designated amplifier to either ×10 or ×1000.

M15_SetGainRange creates and writes a command to the system. The system responds with appropriate message and the function outputs an error message if any.

input = setting (decimal)

= amplifiernum (2-byte sequence from 0 to 32 in ASCII encoded hexadecimal, “00” to                       “20”)

command = [‘R’, amplifiernum, parameter1]

M15_SerialWrite(command)

response = M15_SerialRead

[code, errormsg] = M15_StatusCode(response)

M15_VerifyCode(code,errormsg)

output = none

Parameter1 should be sent in ASCII-encoded hexadecimal, representing the following settings:

Table 2.3.5 Gain Range Setting

| gain range | setting | parameter1 |
| --- | --- | --- |
| ×1000 | 1000 | 0 |
| ×10 | 10 | 1 |

function M15_SetGain(setting, amplifiernum)

This command changes the gain setting for the designated amplifier to the setting represented by the input. The overall gain for the channel is the product of the gain range and the gain setting.

M15_SetGain creates and writes a command to the system. The system responds with appropriate message and the function outputs an error message if any.

input = setting(decimal)

= amplifiernum (2-byte sequence from 0 to 32 in ASCII encoded hexadecimal, “00” to “20”)

command = [‘G’, amplifiernum, parameter1]

M15_SerialWrite(command)

response = M15_SerialRead

[code, errormsg] = M15_StatusCode(response)

M15_VerifyCode(code,errormsg)

output = none

Parameter1 should be sent in ASCII-encoded hexadecimal, representing the following settings:

Table 2.3.6 Gain Setting

| setting | parameter1 |
| --- | --- |
| 5 | 0 |
| 10 | 1 |
| 20 | 2 |
| 50 | 3 |
| 100 | 4 |
| 200 | 5 |

function M15_SaveDefaults

This command saves the current settings to the non-volatile memory so that the amplifiers are initialized to these stored settings on power up.

M15_SaveDefaults creates and writes a command to the system. The system responds with appropriate message and the function outputs an error message if any.

input = none

command = [‘Z’]

M15_SerialWrite(command)

response = M15_SerialRead

[code, errormsg] = M15_StatusCode(response)

M15_VerifyCode(code,errormsg)

output = none